Orifice Flow-meter System

The designed orifice flow-meter system is able to measure flow rates up to 6 GPM. The system consists of a plate with three different diameter orifices, each one being for a different flow range.

| Orifice # | Flow Range (GPM) | Orifice Diameter (in) |
| --- | --- | --- |
| 1 | 0 – 1 | 0.125 |
| 2 | 1 – 3 | 0.202 |
| 3 | 3 – 6 | 0.278 |

When using any particular flow range, the user must make sure to open the ball valves before and after that orifice and close all the other valves.

Special Notes:

The PVC pipe used is a ¾ in pipe, but when the I.D. was measured it was found to be 0.798 in. Maximum pressure for each orifice will be 20 PSI (555 in. of H20). Once the pressure reading on display exceeds 555 in. of H20, it is best to switch to the next larger size orifice. When using orifice #3, the control valve can be opened all the way without worrying about the pressure reading.

Simplified Equations:

| Flow Range (GPM) | Simplified Equation |
| --- | --- |
| 0 – 1 |  |
| 1 – 3 |  |
| 3 – 6 |  |

These equations are a simplified form of

In these equations Qactual  is in GPM (orifice flow-meter volumetric flow rate). All the user has to do is take the reading in (in. of H20) from the digital display and plug it into the corresponding equation. The ABB K-flow flow-meter gives readings in Kg/min, in order to convert that to GPM, the user must divide the reading by 3.785. Readings from the two flow meters can now be compared.

Important:

Before using the orifice flow meter system:

- Make sure there is enough water in the water circuit tank.

- Remove air bubbles from the whole cycle.

K-Values:

| Orifice # | Corresponding K-Value |
| --- | --- |
| 1 | 0.757 |
| 2 | 0.8229 |
| 3 | 0.8068 |

The Flow Coefficient is found using the following equations:

where Qideal  is the volumetric flow rate without losses.

where E is the velocity of approach factor.

where C is the discharge coefficient.